## Sheet1
Hazard Analysis for Biaxial Tensile Tester（Team12）
Part/Process | Potential Failures | Potential Effects of Failure | Severity | Potential Causes | Probability | Current Design Controls | Detection | RPN | Potential Actions
Rake Subsystem | Rake breaks | Tests can not be run or data will be inaccurate | 7 | Rake is too thin | 3 | tolerance stack up/specifications | 3 | 63 | Make rakes thicker
|  |  |  | Rake is bent (besides initial 90 degree bend) | 3 | tolerance stack up/specifications | 4 | 84
| Rake is too sharp | User injury (finger pricked, scratched, etc.) | 9 | Rake is sharpened more than it should be during manufacturing | 4 | Users should wear gloves and if possible use tweezers to connect samples to the rakes. We will be testing each individual rake in terms of sharpness before installation. | 2 | 72
Breadboard and Wires | Loose or disconnected wires | Some actuators or sensors not running properly | 7 | Poor set-up | 3 | put wires firmly in place | 1 | 21 | solder
| Wires get wet | circuit will not run or calculate meaurements | 7 | Exposed wires | 2 | Keep electronics away from water | 1 | 14 | Water-proof covering
Arduino and Motor Driver | Overheats | Will stop running, replace parts | 7 | Too much power/left on too long | 5 | Use low power circuits | 2 | 70
|  | hurt the users |  |  |  | User should wear gloves
Actuator | actuator breaks | the system will not be able to be used | 7 | the spring breaks, poor maintenance | 1 | use springs with higher tolerance | 1 | 7
| motor gets fried | too much power to the actuator | 7 | too much power to the system | 2 | Use low power circuits | 2 | 28
| jams | actuator can't move freely | 9 | someones hair or specimen gets in there | 3 | long hair has to be in a pony tail | 1 | 27
Force sensor | Connectors being damaged | Innaccurate return values | 7 | too much force applied | 3 | sensor can withstand 100N | 1 | 21
Need replacement sensors
Base&Block | cracks or breaks | Affect the measurement, and the accuracy of the system | 6 | holes drilled without pre-drilled holes, device is dropped, water is leaked onto the device | 1 | predrill holes and and do not let water sit on the device | 2 | 12 | Water-proof base and any other wood parts.